TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

TF03 Long-Distance LiDAR

Rs485/RS232

Product Manual V1.3.2

Page1

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Specified Product

Product model: TF03 UART/CAN

Product Name: Long-Distance LiDAR

Manufacturer

Company name: Benewake (Beijing) Co., Ltd.

Address: NO.28 Xinxi Road, Haidian District, Beijing, PRC

Copyright

The Copyright of this document is protected. All the rights involved herein belong to Benewake  (Beijing) Co., Ltd. Any copy activity of this document, no matter in whole or in part, should be in  conformity of the Copyright Law. The actives of modification, omission or translation of this  document are not allowed unless a written permission from Benewake (Beijing) Co., Ltd.

All rights reserved © Benewake (Beijing) Co., Ltd.

Page2

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Foreword

Dear users:

Thanks for choosing Benewake products, and it’s our pleasure to help you to solve any technical  question.

For the purpose of offering a better operation experience to you, we hereby write this manual for  an easier and simpler operation of our product, hoping to better solve the common problems you  maybe meet. Please contact us if you have any questions.( bw@benewake.com)

This operation manual covers the product operation introduction and common problem solutions,  but it is really hard to cover all the problems you maybe meet. So if you have any further questions  or problems, please feel free to consult our technical support service (support@benewake.com).  We will do our best to solve any problem related to the product. If you have any other good advice  or suggestions, welcome to visit our official website and offer us your feedback there  (http://en.benewake.com/support), and we are looking forwards to your participation.

We are Benewake who dedicated to making the best “Robotic Eyes” worldwide!

Page3

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Catalog

1 ATTENTIONS ............................................................................................................................................ 6 1.1 ABOUT THIS DOCUMENT............................................................................................................................. 6 1.2 INSTALLATION AND MAINTENANCE................................................................................................................ 6 1.3 CONDITIONS WITH POTENTIAL PRODUCT FAILURE............................................................................................ 6

2 INTRODUCTION ....................................................................................................................................... 7 2.1 APPEARANCE OVERVIEW ............................................................................................................................ 7 2.2 DIMENSIONS ............................................................................................................................................ 7 2.3 PRINCIPLE OF MEASUREMENT...................................................................................................................... 7 2.4 DETECTION ANGLE DESCRIPTIONS ................................................................................................................. 8

3 PHYSICAL INTERFACE ............................................................................................................................... 9 3.1 WIRING................................................................................................................................................... 9 3.2 ELECTRICAL CHARACTERISTICS.................................................................................................................... 10

4 HARDWARE AND COMMUNICATION PROTOCOL ................................................................................... 11 4.1 HARDWARE PROTOCOL ............................................................................................................................. 11 4.2 COMMUNICATION INTERFACE..................................................................................................................... 11

5 CUSTOM CONFIGURATION..................................................................................................................... 12 5.1 PROTOCOL DESCRIPTION........................................................................................................................... 12 5.2 INSTRUCTION PROTOCOLS ......................................................................................................................... 13 5.3 INSTRUCTIONS OF COMMAND EDITING......................................................................................................... 15

6 MODBUS DESCRIPTION.......................................................................................................................... 15 6.1 PROTOCOL DESCRIPTION........................................................................................................................... 15 6.2 FUNCTION CODE...................................................................................................................................... 16

Page4

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

6.3 REGISTER ADDRESS DESCRIPTION................................................................................................................ 16 6.4 INSTRUCTION PROTOCOLS ......................................................................................................................... 17

7 QUICK TEST PROCEDURE........................................................................................................................ 18 7.1 REQUIRED TOOLS OF PRODUCT TEST........................................................................................................... 18 7.2 TEST PROCEDURE..................................................................................................................................... 18 7.3 TEST EXAMPLE........................................................................................................................................ 20 7.3.1 TEST RANGE WITH DIFFERENT REFLECTIVITY................................................................................................ 20 7.3.2 ACCURACY OF DIFFERENT MATERIALS........................................................................................................ 20

8 CONFIGURATION SOFTWARE OF TF03.................................................................................................... 21 9 FAILURE REASONS AND TROUBLESHOOTING......................................................................................... 22 10 FAQ...................................................................................................................................................... 23 ATTACHMENT 1: REFLECTIVITY OF DIFFERENT MATERIALS .......................................................................... 24

Page5

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

1 Attentions

1.1 About this Document

• This Manual provides information necessary for the use of this product.

• Please read this Manual carefully before using this product and make sure that you have fully  understood its contents.

1.2 Installation and maintenance

• This product can only be maintained by qualified professionals and only the original spare  parts can be used to ensure its performance and safety.

• The working temperature of the product is -25°C~60°C; please do not use it beyond this  temperature range, so as to avoid risks.

• The storage temperature of the product is -40°C~85°C; please do not store it beyond this  temperature range, so as to avoid risks.

• Do not open its enclosure for assembly or maintenance beyond this Manual; otherwise, it will  affect the product performance.

• Do not twist the cable forcefully, so as to avoid damage to product.

• The product can’t be aimed directly to the sun or another TF03, so as to avoid damage the  detector by strong light. If there is such an application, please contact our technician.

1.3 Conditions with Potential Product Failure

• When the product transmitter and receiver lens are covered by dirt, there will be a risk of  failures. Please keep the lens clean.

• The product will have a risk of failure when immersed completely in water. Do not use it  underwater.

• When detecting objects with high reflectivity, such as mirrors and smooth tiles, the product  may have a high risk of failures.

Page6

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

2 Introduction

2.1 Appearance Overview

Figure 1 Appearance

2.2 Dimensions

The LiDAR mounting holes are applicable to M3 screws.

Figure 2 Dimension (Left 1: top; Left 2: bottom; Left 3: front; Unit: mm)

Note: The length of the screw entering the enclosure shall be no more than 3.5mm. 2.3 Principle of Measurement

TF03 is a single point LiDAR, which is based on Pulse Time of Flight (PTOF). It adopts an  incoherent energy receiving mode, and the detection is mainly based on Pulse counting.

TF03 emits a narrow pulse laser, which is collimated by the transmitting lens to form a collimated

Page7

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

light, which enters the receiving system after being reflected by the measured target and is  focused on the APD detector by the receiving lens. The time between the transmitted signal and  the received signal is calculated through the circuit amplification and filtering, and the distance  between TF03 and the measured target can be calculated through the speed of light.

Laser

Pulse timing

Module

Figure 3 Principle of measurement

2.4 Detection angle descriptions

TF03 has a 0.5-degree detection angle and rectangular light spot, see Figure 4 for simulated  diagrams of the light spot. Therefore, at different distances, the spot size, namely detecting range,  is different as shown in Table 1.

Note: The side length of common objects detected should be greater than that of the detection  range of TF03; When the side length of the detected object is less than that of the detection range,  the LiDAR effective range will be reduced

Figure 4 Detection angle of TF03

Page8

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Table 1 Spot simulated diagrams and sizes at different distances

| Spot 20m away | Size：20cm*6cm | Spot 50m away | Size：50cm*14cm |
| --- | --- | --- | --- |
| Spot 100m away | Size：100cm*28cm | Spot 150m away | Size：150cm*42cm |

Note: In these pictures, color represents the light intensity distribution, the red is the strongest,  the green is weak, the blue is the background. The spot size is the length and width of the green  in the picture.

3 Physical Interface

New version of TF03-100/180 is released since August 2020, the physical interface is adjusted  as follows:

⚫ Bold cable, the size of single wire is changed from AWG30 to AWG 26 ⚫ Remove a wire, the amount of wire is 6 now, color is shown in Figure 5 ⚫ Add overvoltage and polarity protection, shown as Table 3

3.1 Wiring

Terminal model: MH1.25-7P-W/B, size of single wire is AWG26, diameter of single wire is

Page9

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

0.404mm, cross-sectional area is 0.129mm2.

Figure 5 New Line Sequence of TF03

Table 2 Pin functions and connection instructions of TF03

| No. | Color | Pin | Function |
| --- | --- | --- | --- |
| 1 | Red | VCC | Power supply |
| 2 | White | RS485-B/RS232-RX | RS485-B/RS232 receive |
| 3 | Green | RS485-A/RS232-TX | RS485-A/RS232 transport |
| 4 | N/A | N/A | N/A |
| 5 | Blue | UART_RX | UART receive(debug1) |
| 6 | Brown | UART_TX | UART transport(debug) |
| 7 | Black | GND | Ground |

Note: If UART is used, the TTL_RXD and TTL_TXD must be connected and not allowed to be left  floating, so as not to be affected by electromagnetic interference.

3.2 Electrical Characteristics

New version has overvoltage and polarity protection.

Table 3 Main electrical parameters

| Parameter | Typical value |
| --- | --- |
| Supply voltage | 5V~24V |
| Average current | ≤150mA |

1 The UART interface of TF03-100 RS485 version is debugging interface. It cannot be used to read detection data.

Page10

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

| Average power consumption | ≤1W |
| --- | --- |
| Overvoltage protection | 300V |
| Polarity protection | 200V |

4 Hardware and communication protocol

The industrial version of TF03 supports two communication interfaces, RS485 and RS232. The  default interface is RS485. These two interfaces cannot work simultaneously. The communication interface can be switched by command.

4.1 Hardware protocol

TF03 supports two communication interfaces, RS485 and RS232. See the hardware protocol  parameters below.

Table 4 Hardware protocol of TF03

| Item | Parameter |
| --- | --- |
| Communication interface | RS485/RS232 |
| Baud rate | 115200 |
| Data bit | 8 |
| Stop bit | 1 |
| Parity | None |

4.2 Communication interface

The output data (hexadecimal numbers) of TF03 is shown in

Page11

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Table 5. Each data frame consists of 9 bytes and the data contains the distance and signal  strength2information.

2 New version of TF03 adds strength information. Strength value is between 0 and 3500, threshold of strength is 40, when strength is lower than 40, dist will output maximum value. When strength is between 40 and 1200, dist is more reliable. When  there is a high reflectivity object, strength will be over 1500.

Page12

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Table 5 TF03 UART data format

| Data bit | Definition | Description |
| --- | --- | --- |
| Byte0 | Frame header | 0x59 |
| Byte1 | Frame header | 0x59 |
| Byte2 | DIST_L | DIST low 8-bits |
| Byte3 | DIST_H | DIST high 8-bits |
| Byte4 | Strength_L | Signal strength low 8-bits |
| Byte5 | Strength_H | Signal strength high 8-bits |
| Byte6 | Reserved bit | / |
| Byte7 | Reserved bit | / |
| Byte8 | Checksum | Low 8 bits of Checksum bit. Checksum = Byte0 + Byte2+…+Byte7 |

5 Custom configuration

5.1 Protocol description

To meet the need of different customers, parameters can be set by yourselves. Parameters, such  as data format, frame rate could be changed by sending command. Parameter will be stored in  flash after configured successfully and customers don’t need to configure again when restart.

Please change the parameter according to certain demands and do not frequently try irrelevant  instructions. Please configure the product according to the demands of the manual and never  send unstated command.

Table 6 Description of TF03 instruction protocol

| Byte | Definition | Description |
| --- | --- | --- |
| Byte0 | Head | Fixed to 0x5A |
| Byte1 | Len | The length of the entire instruction frame (unit: Byte) |
| Byte2 | ID | Identifies the function of each instruction |
| Byte3~ByteN-2 | Payload | Different meanings and lengths in different ID instruction frames |
| ByteN-1 | Check sum | the lower 8 bits of the Len-1-byte data |

Page13

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

5.2 Instruction protocols

Table 7 List of TF03 instruction protocols

| Function | Command | Response | Remark | Default setting |
| --- | --- | --- | --- | --- |
| Obtain firmware version | 5A 04 01 5F | 5A 07 01 V1 V2 V3 SU | The version number V3.V2.V1 SU is check sum | / |
| System reset | 5A 04 02 60 | Succeed: 5A 05 02 00 61 Fail: no response over 1s | / | / |
| Frame rate3 | 5A 06 03 LL HH SU | Succeed: 5A 06 03 LL HH SU Fail: no response over 1s | / | 100fps |
| Output control | Enable: 5A 05 07 01 67 Disable: 5A 05 07 00 66 | Succeed: same as downward command Fail: no response over 1s | / | Enable |
| Trigger command | 5A 04 04 62 | Data frame | / | / |
| Output format | 5A 05 05 LL SU | Succeed: 5A 05 05 LL SU Fail: no response over 1s | LL: format shown as below, 00: ASCII (Reserved) 01: Binary output | Binary output |
| Modify baud rate4 | 5A 08 06 H1 H2 H3 H4 SU | Succeed: 5A 08 06 H1 H2 H3 H4 SU | See chapter 4.3.3 | 115200 |

3 New version of TF03 can only support (1~9) *10N, N values for 0, 1, 2, 3. If other value is set, TF03 will set it to 100Hz. If  10000Hz is necessary, please contact our technician.

4 Baud rate of UART can be set to 9600, 14400, 19200, 38400, 56000, 57600, 115200, 128000, 230400, 256000, 460800,

Page14

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

|  |  | Fail: no response over 1s |  |  |
| --- | --- | --- | --- | --- |
| Enable checksum | Enable: 5A 05 08 01 68 Disable: 5A 05 08 00 67 | Succeed: same as command Fail: no response over 1s | / | Enable |
| Restore factory settings | 5A 04 10 6E | Succeed: 5A 05 10 00 6F Fail: 5A 05 10 ER SU | Fail (when ER is not 0) | / |
| Save settings | 5A 04 11 6F | Succeed: 5A 05 11 00 70 Fail: 5A 05 11 ER SU | same as above | / |
| Over range threshold value | 5A 06 4F LL HH SU | Succeed: 5A 05 4F 00 AE Fail: no response over 1s | Unit: cm See chapter 4.3.3 | 18000 |
| Communication interface | UART: 5A 05 45 01 A5 CAN: 5A 05 45 02 A6 | Succeed: 5A 05 45 00 A4 Fail: no response over 1s | / | UART |
| Rain-fog algorithm | Enable: 5A 05 64 00 C3 Disable: 5A 05 64 01 C4 | Succeed: 5A 05 64 00 C3 Fail: no response over 1s | / | Enable |
| Offset configuration5 | 5A 06 69 LL HH SU | Succeed: 5A 05 69 00 C8 Fail: no response over 1s | Unit: cm See chapter 4.3.3 | 0 |
| Switch communication interface | 5A 05 45 CI SU | Succeed: 5A 05 45 00 A4 Fail: no response over 1s | CI: communication interface 0x01: RS232 0x03: RS485 | RS458 |

500000, 512000, 600000, 750000, 921600 and 1000000, if other value is set, TF03 will set it to 115200.

5“Offset configuration” can be used for secondary calibration of distance, for example, when dist is 195cm and you want LiDAR  outputs 200cm, you can set offset value to 5cm. This function is supported from v1.11.3 firmware version.

Page15

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

| Enable Modbus | 5A 05 6F 00 CE | Succeed: 5A 05 6F 00 CE Fail: no response over 1s | Restart to take effect | Disable |
| --- | --- | --- | --- | --- |
| Set Modbus address | 5A 05 70 AD SU | Succeed: 5A 05 70 00 CF Fail: no response over 1s | AD: Modbus communication address. Effective range: [1,247] | 0x01 |

5.3 Instructions of command editing

To send certain command to TF03, these steps should be followed:

⚫ Change value of parameter to HEX

⚫ Put the HEX value to certain command

⚫ Calculate the checksum, put low 8-bits into command

For example, changing the baud rate to 460800, first, changing 460800 to HEX, we can get 0x00  07 08 00, then we can get this command-5A 08 06 00 08 07 00 77.

6 Modbus description

TF03 Modbus adopts a 2-wire interface. For the line sequence, see the physical interface  description-line sequence description; for the check and baud rate setting, see the communication  protocol and data format. For custom parameter configuration, you need to switch to serial port  mode to set the baud rate.

The TF03 RS485 interface is in half-duplex mode. Based on reliability considerations, it is not  recommended to use a baud rate above 115200 for communication.

6.1 Protocol description

When Modbus protocol is enabled, the Modbus reading distance command format:

Table 8 Distance reading command of TF03 RS485 Modbus

| Address field | Function code | Register address |  | Number of bytes read |  | CRC_low | CRC_high |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 01 | 03 | 00 | 00 | 00 | 01 | xx | xx |

The data frame returned by TF03 is as follows:

Page16

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

| Address field | Function code | Data length | Dist_high | Dist_low | CRC_low | CRC_high |
| --- | --- | --- | --- | --- | --- | --- |
| 01 | 03 | 02 | xx | xx | xx | xx |

6.2 Function code

The functions supported by the Modbus protocol are relatively simple, only two function codes  are open

Table 9 Function code of TF03's Modbus

| Function code | description |
| --- | --- |
| 03 | Read register |
| 06 | Write register |

6.3 Register address description

Table 10 Register address list of Modbus protocol (Function code: 0x03)

| Register address | Definition | Description |
| --- | --- | --- |
| 00 00 | Dist | Distance |
| 00 01 | Strength | Signal strength |
| 00 03 | Timestamp (16 high-order bits) | Unit: ms |
| 00 04 | Timestamp (16 low-order bits) | Unit: ms |
| 00 06 | High16bit of software version | 00 + major version number |
| 00 07 | Low16bit of software version | Minor version number + revised version number |

Table 11 Register address list of Modbus protocol (Function code: 0x06)

| Register address | Definition | Description |
| --- | --- | --- |
| 00 80 | Save settings | Write any value to take effect |
| 00 81 | Restart | Instruction value: 0x01 |
| 00 82 | Disable Modbus | Instruction value: 0x01 |
| 00 83 | Baud rate (16 high-order bits) | Restart to take effect |
| 00 84 | Baud rate (16 low-order bits) | Restart to take effect |

Page17

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

| 00 85 | Slave ID | Set TF03’s ID. Restart to take effect |
| --- | --- | --- |
| 00 86 | fps | Restart to take effect |
| 00 87 | Working mode | 0x00: Continuously detection mode 0x01: triggering mode |
| 00 89 | Restore default | Write any value to take effect |

6.4 Instruction protocols

The default address field is 0x01.

Table 12 Instruction protocols of Modbus

| Category | Function | Instruction | Response | Description |
| --- | --- | --- | --- | --- |
| UART instructio ns | Enable Modbus | 5A 05 15 01 75 | 5A 05 15 01 75 | Save settings and restart to take effect. |
|  | Save settings | 5A 04 11 6F | 5A 05 11 00 70 |  |
| Modbus instructio ns | Read distance | 01 03 00 00 00 01 84 0A | 01 03 02 DH DL CL CH | DH: 8 high-order bits of distance DL: 8 low-order bits of distance CH: 8 high-order bits of CRC CL: 8 low-order bits of CRC |
|  | Read distance and signal strength | 01 03 00 00 00 02 C4 0B | 01 03 04 DH DL SH SL CL CH | DH: 8 high-order bits of distance DL: 8 low-order bits of distance SH: 8 high-order bits of signal strength SL: 8 low-order bits of signal strength |
|  | Read software version | 01 03 00 06 00 02 24 0A | 01 03 04 00 VM VS VC CL CH | The software version is VM.VS.VC |

Page18

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

|  | Set Baud rate | 01 06 00 83 BH1 BH2 CL CH 01 06 00 84 BL1 BL2 CL CH | 01 06 00 83 BH1 BH2 CL CH 01 06 00 84 BL1 BL2 CL CH | For example, set baud rate to 9600 (0x00002580) BH1=00 BH2=00 CL=78 CH=22 BL1=25 BL2=80 CL=D2 CH=D3 |
| --- | --- | --- | --- | --- |
|  | Set Slave ID | 01 06 00 85 IH IL CL CH | 01 06 00 84 IH IL CL CH | Set slave ID to 0x02，IH=00 IL=02 CL=19 CH=E2 |
|  | Set frame rate | 01 06 00 86 FH FL CL CH | 01 06 00 86 FH FL CL CH | Set frame rate to 100 (0x0064) ，FH=00 FL=64 CL=69 CH=C8 |
|  | Save settings | 01 06 00 80 00 00 88 22 | 01 06 00 80 00 00 88 22 | Restart to take effect |
|  | Disable Modbus | 01 06 00 82 00 01 E8 22 | 01 06 00 82 00 01 E8 22 | Save settings and restart to take effect |

7 Quick Test Procedure

7.1 Required Tools of Product Test

Note: the product package contains only TF03 and factory certificate, other accessories are  prepared by yourself. If you need RS485-USB convert board, please contact sales or technical  support. The UART version needs a RS485-USB board when connected with computer. Tools  and method are shown as below.

Table 13 Tools of test

| TF03(UART) | RS485-USB Board | USB Cable | PC | TF display application |
| --- | --- | --- | --- | --- |

7.2 Test procedure

1）Download the TF display application

Page19

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Please download the TF display application from http://en.benewake.com/supportNote:  Please close the antivirus software before unpacking the TF display application, otherwise the  files will be deleted as viruses. The TF display application currently only supports running on the  Windows system. See Attachment 1: Instruction of TF Series PC Display Software.

2）Connecting

Figure 6 Diagram of connection

As Figure 6 shown, connect 『TF03』 and 『RS485-USB board』 and 『USB cable』, ensure  no looseness, then connect 『USB cable』with 『computer』.

3）TF display software connection and reading

C

A

B

D

Figure 7 The interface of TF display software

As shown in Figure 7, test steps are shown as below:

➢ Open the TF display application

Page20

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

➢ Select product model『TF03』 in area A

➢ Select serial port number in area B.

➢ Click『CONNECT』to establish the communication.

➢ After connected Successfully, continuous data curve will be shown in the area C 『TIME  LINE CHART』 and area D 『REAL TIME DATA』 will display the current distance (Unit:  cm) and frame rate (Effective Points).

7.3 Test Example

7.3.1 Test Range with Different Reflectivity

As shown in Figure 8, red and blue curves represent the relationship between range performance  of TF03-180 and reflectivity of the target under different ambient light situation.

Figure 8 The relationship between range performance of TF03-180 and reflectivity

7.3.2 Accuracy of Different Materials

The accuracy of TF03 at different distances and reflectivity is shown in Figure 9. Two typical  background boards(black and white) are selected. The reflectivity of whiteboard is 90% and  blackboard is 10%.

Page21

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Figure 9 Error curves of TF03 at different distance and reflectivity

(Test environment: 25℃, indoor without ambient light)

8 Configuration Software of TF03

V1.3.2

Figure 10 The interface of configuration software

Page22

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Interface of configuration software is shown above.

1) After TF03 is connected to the computer, select the correct serial port number and baud rate  in Zone C manually and click the 『Connect』 button to establish communication between  the software and TF03.

2) After communication is established successfully, zone B will display the distance and  frequency of data. When no data is received over two seconds, the distance and frequency  display will disappear.

3) Zone D is dedicated to function configuration and its third column lists the buttons for sending  instructions. After clicking such buttons, the configuration software will send commands and

wait for reply from TF03. When no reply is received for a long time, 『No Response』will be  displayed on the right side of the instruction-sending buttons. Please note that in order to  ensure that TF03 reply instructions can be obtained normally, lower the『frame rate』or set

the『output mode』as『instruction trigger』before configuration. If the changed parameters  need to be saved, click 『Save Configuration』before power down TF03. 4) Zone E is dedicated to firmware upgrade, Bootloader can be used to upgrade the firmware. 9 Failure reasons and troubleshooting

1）Under normal operation, TF03 sometimes will output 18000(cm).

Reasons：The actual distance measured is beyond the range of TF03, TF03 will output 18000.

Troubleshooting：Treat 18000 as an exception value. When you receive 18000 from TF03, don't  adopt this data.

2）No data output after TF03 is connected to the TF display application.

Reason Ⅰ：The computer uses and operating system other than Windows.

Troubleshooting：Currently, the TF display application can only support the Windows operating  system. Use a computer with the Windows system installed.

Reason Ⅱ：Poor connection between TF03 and computer.

Troubleshooting：Please confirm that TF03 is correctly and reliably connected to the computer,  and ensure that the cables and adapter board work properly.

3）TF03 has a poor range performance in some cases.

Page23

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Reason：TF03 bases on the time-of-flight (TOF) principle, to achieve ranging by calculating the  round-trip time of flight of laser pulses. The intensity of ambient light or the reflectivity of objects  under testing is different, and thus will increase the intensity of ambient noise or return light  intensity. Therefore, the ranging performance of TF03 varies with different ambient light intensities  and reflectivity.

Troubleshooting：Under different ambient light intensity and reflectivity of objects, having different  range performance is a normal phenomenon and does not affect the accuracy or repeatability.

10 FAQ

Q1: Can the FOV (spot) of TF03 be increased or decreased?

A1: Generally, this is a customized demand. You need to contact our sale colleague for further  details. FOV is determined at the beginning of product design. Determination of FOV is also  related to the optical system and the product structure. Therefore, it cannot be easily changed,  and needs to be customized.

Q2: Can I change the frequency of data output? Can TF03 output switching values?

A2: Yes, customization of parameter configuration and adjustment of output modes have been  enabled for TF03. However, switching values have to be customized according to requirement.  You may consult our sales colleague or technical support for further details.

Q3: Can TF03 work normally under raining, snowing and fog?

A3: TF03 has a firmware version which support this function. Contact our sale colleague or  technical support for firmware upgrade. The algorithm is turned on by default in V1.11.15.

Page24

TF03 RS485/RS232 Product Manual V1.3.2 Benewake(Beijing)Co., Ltd

Attachment 1: Reflectivity of Different Materials

The reflectivity of different materials is listed below, ranging from low to high. According to the test  target and the corresponding reflectivity, we can measure whether the range of TF03 and other  parameters meet the requirements.

| No. Materials Reflectivity |  |  |
| --- | --- | --- |
| 1 | black foam rubber | 2.4% |
| 2 | black cloth | 3% |
| 3 | black rubber | 4% |
| 4 | Coal (varies from coal to coal) | 4~8% |
| 5 | Black car paint | 5% |
| 6 | Black paper | 10% |
| 7 | opaque black plastic | 14% |
| 8 | Clean rough board | 20% |
| 9 | newspapers | 55% |
| 10 | translucent plastic bottles | 62% |
| 11 | packing case cardboard | 68% |
| 12 | Clean pine | 70% |
| 13 | opaque white plastic | 87% |
| 14 | white card | 90% |
| 15 | Kodak standard whiteboard | 100% |
| 16 | Unpolished white metal surface | 130% |
| 17 | Shiny light metal surface | 150% |
| 18 | stainless steel | 200% |
| 19 | Reflective board, reflective adhesive tape | >300% |

ContactHeadquarters:

400-880-9610

Technical support:

bw@benewake.com support@benewake.com

Page25